## General
Liens list from the VIR LAMO/XMO1/XMO2 L1a and L1b review
Aug. 1, 2017
Attending:
Richard Chen | PDS EN
Bethany Ehlmann | External reviewer
Marco Giardino | Dawn VIR team
Steve Joy | Dawn DSC
Ludmilla Kolokolova | PDS SBN
Joe Mafi | Dawn DSC
Ralph Milliken | External reviewer
Bea Mueller | PDS SBN
Carol Neese | PDS SBN
Eric Palmer | PDS SBN
Carol Raymond | Dawn Deputy PI
Certification:
These data can be provisionally released to the public with the existing documentation IF the ground correction factor is provided.
The information provided for this will be confirmed by Ralph before release.
Follow up and completion:
Reviewers agree to participate by email in checking drafts of the revised documentation and providing feedback.
The completed documents will be included in the upcoming final VIR review.
The deliveries for the final VIR review, including the revised documentation, are due by Aug. 31.
Since the VIR team will be unavailable for the month of August, it is expected this date will slip somewhat.

## Cal doc
Reviewer | Lien | SBN addition | VIR comment | Reviewer comment (BE) | Resolution comments | Completion date
BE | There are some crucial issues with the VIR_CALIBRATION document as well as the VIR_LAMP_DESCRIPTION document, in particular the consistency of the description of the instrument response function (IRF) derivation relative to the what is described in the SIS and CAT files. In particular, the V2 IRF derivation is not sufficiently described. Indeed, the calibration document implies a completely on-the-ground IRF derivation while the SIS and the L1B VIR_VIS_RDR_CERES_DS.CAT files imply onboard information was used. Need to explicitly describe the basis of the DAWN_VIR_*_RESP_V2.TXT information and how onboard lamp data were used. | This lien is highlighted because of its importance. | Will do. | Great
BE | A second crucial issue is that there should also be a systematic listing of known artifacts in either the calibration document or the L1B VIR_VIS_RDR_CERES_DS.CAT. It can be one or the other and not both, but a cross reference needs to be explicitly called out | This lien is highlighted because of its importance. This information will be included all in one place in the calibration document, and be referred to in the data set cat. | Such information is provided in the quality qubes described in the SIS. We can consider adding a paragraph to cal doc to summarize all known artifacts and refer to the CAT file for an extensive description. | All artifacts need to be listed in a single place, systematically. The calibration document is the logical place to do this, but if the team instead wishes to make more comprehensive the current list in the CAT that is okay (with a sentence in the Cal Doc explicitly refering the reader to a list there). The quality qubes are not sufficent as some verbal explanation of the causes, severity of, potential for fixes for the artifact is needed
BE | Section 3.1 Does the spectral dispersion equation provide the wavelength position? If so, at first or second use of the term "spectral dispersion" please state this explicitly as it was not one of the calibration parameters defined on p3. (or definite on p3, section 3 intro) |  | With Bethany's clarification, Marco says this lien is now clear. | The term "spectral dispersion" is used several times and the means of computing it given. But the term is not defined. I believe it is equivalent to what is defined on p3 as "Sample Spectral Wavelength". So need to define "spectral dispersion" on p3 or use the term provided on p3 for the quantity.
BE | p8 I don't quite understand this part of a very important sentence: "but the knowledge on the geometric factor (constant and uniform for each spectral channel) is not known." Please rephrase. Does it mean "but the absolute radiance cannot be known because there is a different geometric factor for each channel"? |  | Will do. | Great
BE | p9 Section 7. Please differentiate and describe the measurement procedure to obtain electronic offset, background, dark current, acquisition of lamp data |  | Will do. | Great
BE | p10 If the original data are shifted two pixels, I don't understand how the detilt algorithm is used to produce an array of the original size, as stated. It seems to me that the array must be 2 pixels smaller in the spatial dimension, because data that would correspond to 2 pixels was not received by any detector element. Please clarify why the original dimensions hold while dealing with a 2 pixel shift of the light field. |  | Will do. | Great
BE | p12 Typo? The first term defined under the equation is not equivalent to the lefthand side of the equation immediately proceeding. Should it be S(λ(b),s,l)? If not, please clarify how to solve te equation given |  | Will do. Probably it should be sufficient to change x with s and y with l across the page to improve the readibility. | I will leave you to determine what is correct.
BE | Section 8: please specify explicitly. Does detilt occur before all of the steps involving the IRF? |  | Will do. | Great
BE | Section 8: Please parse the text, splitting into clearer sentences as it is not clear what the current issues are. "The actual ITF is also currently under improvement: calibrated values in the spectral range [2.818μm - 3.272μm] are still under verification, this is the reason why these values have been put to null into the ITF only for the acquisitions performed during the mission campaigns VSH and VH2 : during these two campaigns, a major outer contamination is taken as responsible for the artefacts affecting the spectral range around 3 micron." Is the "a major outer contamination" material on some part of the instrument that was later removed? Is the ground ITF under investigation or just the effective instrument response during those two mission periods (or both)? |  | The explanation about the absence of contamination during other campaigns will be documented. The team will also deliver the artifact removed version of the Vesta dataset. where the contamination will be partially removed. | Great. Please just make sure this is clearly written in this section
BE | Section 8: Please comment explicitly on whether the IRF provided is that derived from ground calibration or whether any of the section 7 measurements have been used to create the current IRF. The 'VIR_VIS_RDR_CERES_DS.CAT' (DWNCLVIR_V1B) seems to imply that either the IRF is not from the ground calibration OR that the user interested in analysis of the 2.5-3.5um region crucial for hydroxyls, must instead use the DeSanctis+ 2012 method rather than the PDS data. This is incredibily important scientifically and must be discussed in the calibration document. Should update to include any differences between dark Ceres and brighter Vesta, if any. Please, if the ITF used to generate the 1B data provided is from the lamp, please step through explicitly in section 7 the steps for how the onboard information was used to recalibrate the IRF with the onboard data. If not, state explicitly under what circumstances the user might want to start with L1A data and do the DeSanctis+ 2012 procedure rather than use the L1B data. | Clarify the text to avoid this misunderstanding. This lien is considered to be very important. | The alternative method must be used for that region when using the 1B dataset and ITF v1. With ITF v2 the alternative method is no more applicable and must be avoided. The artifact removed dataset will solve this region too. The sections 7 and 8 will be improved accordingly. | Great. Please just make sure this is clearly written in this section as this is very crucial
BE | ***UPDATE after reading SIS p18, last parag: it appears a new ITF was made. In this case the procedures should be described explicitly, step by step.
| Section 8: I think it is essential and a great service to the community to include a comprehensive list the spectral positions of known artifacts or potential issues under certain observing scenarios. This will prevent the community from interpreting, e.g. filter boundaries, the 2.5 micrometer region (mentioned elsewhere in the docs as challenging), and other regions. (from my experience with CRISM data, this is important, as I still review a smattering of papers per year that try to interpret our filter boundaries as real data because they didn't read the calibration paper. So best to put the artifacts with the data processing instructions and update as needed with new cal) Right now, the discussion is spread among many documents, with most buried in 'VIR_VIS_RDR_CERES_DS.CAT'. At the very least, the calibration document should explicitly refer to the description of artifacts in'Processing' and 'Data Coverage and Quality' 'VIR_VIS_RDR_CERES_DS.CAT' and ensure that the listing in 'Data Coverage and Quality' is clear comprehensive (right now it is not as some of the important information on artifacts is found instead in 'Processing'). | This information is to be included in a separate section of the calibration document, and should be referred to in the data set cat. If possible, include the Carrozo et al. paper in the documents of the archive. | See #5. An official VIR team documentation about artifacts has been recently published (December 2016). This doc will be a part of the future artifact removed dataset. It could be further investigated the possibility to include it also in the dataset currently under review . | Yes. If this document already exists and is comprehensive, could provide reference with only a brief summary, e.g. a table here with reference to the details in the publication, with full reference provided (also relates to line 5)
RM | The overall calibration documentation and steps for processing the VIR data from DN to radiance to I/F is not nearly as comprehensive as for other similar instruments (e.g., OMEGA, CRISM, M3, etc.). | Ralph will send a more specific list.
RM | Specifically, the keystone correction (detilt algorithm )is not adequately described. | Ralph provided specific questions in past reviews which are not addressed yet. He says he will collect these and send us.
RM/BE | In the detilt algorithm, is the 2 pixel offset a constant offset? | Marco confirms it is a constant 2 pixels spatially, not variable. These details should be described clearly in the documentation of the detilt algorithm. | Marco agreed that the VIR team will add an explanation of this to the document.
RM | In addition, the description of exactly how the wavelengths in the ~3-5 um region were corrected (using onboard lamp) are not fully described, or at least there is not enough detail for the user to be able to recreate this step. | The information is buried in the calinfo.txt and is only an overview with insufficient detail. the whole process of how the second ITF was determined needs to be described in greater detail. | Will do, however, without the polystyrene spectrum the step cannot be completely recreated.
RM | It would be good to include the polystyrene transmission spectrum that was used by the team for this process. This means the table of numbers, not just a plot. | The publication is 15 years old, it may be possible to get it released. Chris Russell has encouraged the team to try to get the spectrum released for inclusion in the archive. | The team is unable to disclose this information, due to legal issues related to the industrial partner policy. However, they agreed to try to get it released.
RM | The team has used a higher level of processed data in many publications, including data that have been filtered for the systematic oscillations in the spectral domain as well as implementation of the ‘ground correction’. These issues are now documented in a peer-reviewed publication and in that publication it is demonstrated that fixes have been implemented. I can understand that the team may release such improved data as Level 1C, but there is no reason that the ground correction vector can’t be released now for PDS users. That is, the team should release the ground correction factor as part of this release since the data are clearly not accurate without it. As such, the VIR data set has limited use without the community having access to this correction factor. I am not requesting that the team release the newly corrected data at this point, simply that they release the ground correction factors such that users can implement it themselves. | Release of the ground correction is very important for the useability of the data. The certification of the data is contingent on the release of this correction factor. | We agree to release this factor also within the current 1B dataset.
BE | p1 Typo, apostrophe needed and plural "detectors['] performance[s]" or change to "the performance of the two detectors" | Please correct these typos and minor errors.
p1 Typo "section 8 explain[s]"..."raw data [into] spectral radiance..."
p4 Typo. As written, not a sentence. Suggest changing to "The calibration setups used to define the spectral properties of VIR [used] a heritage setup developed for the VIRTIS/M aboard Rosetta and the Venus Express missions"
p4 Typo, extra word "Moreover on the optical bench <<<were>>> a spectral shift along the slit is apparent, given that at every illuminated band a different central wavelength is found."
p4 Bottom, Table and p5 Bottom, table formatting problem
p5 Typo. Transmission Method and Diffusion Method sections are both numbered 3.1
p5 Typo. "discussed in advance" should be "discussed previously"
p7 Typo extra word to delete "For the VIS channel the FWHM of the pixel function <<<is>>> ranges"
p7 Please specify what wavelengths are meant by "the two spectral extremes of the range."
p7 typo apostrophe needed "microlamps['] position"
p7 Typo "associate[d] barycenter"
p8 Typo "a diffusive target <<<in>>> [of] Spectralon"
p9 Typo verb "consist[s] in acquistion"
p10 Typo "A dedicated" change to "Dedicated"
p11 Typo wrong word "DARK_ACQUISITION_RATE <<<into>>> [in] the data cube label file."
p11 Typo change adjective to adverb "of two consecutive[ly] acquired"

## Lamp doc
| Lien | SBN comment | VIR comment | Resolution comment | Completion date
| Add an introductory paragraph to the document (or enhance the one already there) to explain that the information is all extracted from the paper, which pertains to the VIRTIS lamp, and that it is provided because the similarity of the VIR lamp to the VIRTIS lamp makes the information relevant. The document should be about the lamp, and information about how the lamp was used to determine the ITF should go in the calibration document. | Steve appears to be stuck doing this one or at least supervising it, although it is rather outside his expertise, He can call on Bethany and on the VIR team for help as needed.
Typo, extraneous words <<>> in sentence: "During the cold tests the lamps were cooled down <<<the lamps>>> to an ambient..."
Typo, "On the right side <<<pyroxene>>>[polystyrene?] features are not present in this band region."
| Fifth, please improve the reproduction quality of the figures as many lines cannot be discerned. And what are the multiple curves in Figure 17 (if this figure remains in the revision)? | Do if reasonably possible.
Remaining reviewer comments below are included in case it makes sense to do any of them. They are not liens.
Reviewer | Additional reviewer comments (not liens)
BE | Please note that it is not necessary to make this document longer. Rather shorter but more clearly oriented to the essential information used in the flight-derived IRF correction is needed.
BE | First, some clearly extraneous information should be removed (filament length details, the argon lamp data and why-an-argon lamp was not chosen, the measurements at times longer than 5min not relevant to flight conditions)
BE | Second, consider stating the objective (given in the last paragraph) earlier: to determine the lamp temperature based on resistance. This is the last paragraph "In flight..." and the equation used to make the correction. Stating this objective earlier would make the discussions of the various calibrations clearer because they would flow directly from this most important aspect.
BE | Third, which equation is used for the flight correction? That on Figure 15, Table III, or Eqn 1? It is not clearly relevant to present the others if they are not used, unless as a point of comparison, in which case a comparison should be made explicitly. Also, Figure 15 'parameter a' may be missing a decimal place.
BE | Fourth, please describe in the revision at an appropriate place what wavelength range is measured. Especially for the IR lamp and after reading section 7 of the calibration document, it is not clear how the wavelength calibration was done for the key range for VIR IR if all the absorptions of the polystyrene are not recorded. Why is the calibration shown only for wavelengths <1200nm when the lamp is used in flight for longer wavelength data out to ~5 micrometers?
BE | The second paragraph of the .txt file reads like the text is about calibration of VIRTIS-M. Need to explicitly describe the relationship to Dawn VIR. Are these two lamps described the lamps actually used in VIR? Are they merely similar? Please rephrase

## SIS
Reveiwer | Reviewer comment | SBN comment | VIR comment | BE comment | Resolution comments | Completion date
BE | p18, last parag. Please state explicitly whether the team believes that the Ceres data after VSH are affected by cold trapped volatiles. That is, please be explicit about which Ceres phases have this uncorrected-for-artificat and which do not. Some brief explanation of how this is known would also be helpful (instrument flats?). |  | Will do | Great
BE | Fig. 4. Where does the detilt described in the calibration document happen in the VIR VIS processing? |  | Will add the detilt step to the diagram | Great
BE | Table 11 should be updated to indicate the V2 versions of the files given, since these are now what is provided in the delivery |  | Will do | Great
BE | Appendix C should also explicitly mention the ENVI and ISIS tutorials in the DOCUMENT folder as this will be very commonly employed |  | Will do | Great
BE | Table 14 seems to mix multiple fonts. May wish to correct. |  | Will do | Great
BE | Fig. 1 appears to have its aspect ratio strangely altered so that all the planetary bodies appear like squished ovals. Fix. |  | Will do | Great

## Data labels
Reviewer | Reviewer comment | SBN comment | VIR comment | Resolution comments | Completion date
The liens in this section will be resolved by DSC, with consultation if necessary with Richard and the VIR team.
RC | I've previously mentioned (and waived?) this? For many data labels, If convenient, please change |  | In our understanding, this was considered already fixed.
TARGET_NAME = "CAL LAMP" to "CAL LAMPS"
RC | DWNC1VIR_I1B/CALIB/DAWN_VIR_IR_RESP_V2.LBL
DWNC1VIR_V1B/CALIB/DAWN_VIR_VIS_RESP_V2.LBL
DWNC2VIR_I1B/CALIB/DAWN_VIR_IR_RESP_V2.LBL
DWNCLVIR_I1B/CALIB/DAWN_VIR_IR_RESP_V2.LBL
DWNCLVIR_V1B/CALIB/DAWN_VIR_VIS_RESP_V2.LBL
- All have
| AXIS_UNIT = ("DIMENSIONLESS","DIMENSIONLESS") |  | Ok to use N/A.
OBJECT = ELEMENT
DESCRIPTION = "DAWN-VIR VIS responsivity in
(DN*m^2*um*sterad)/(W*s)"
and no UNIT within object ELEMENT. That is legal. However, I suggest dropping
AXIS_UNIT since "DIMENSIONLESS" is a new value. Also, would it make sense
to add UNIT within ELEMENT?
RC | DWNC*_I1B/CALIB/*.LBL
- This DATA_SET_ID is currently unknown, so please ensure it is correct
DATA_SET_ID = {"DAWN-A-VIR-3-RDR-IR-VESTA-SPECTRA-V2.0", ...}
RC | DWNC*_V1B/CALIB/*.LBL
- This DATA_SET_ID is currently unknown, so please ensure it is correct
DATA_SET_ID = {"DAWN-A-VIR-3-RDR-VIS-VESTA-SPECTRA-V2.0", ...}
RC | DWNC2VIR_I1A/CATALOG/VIR_VIS_EDR_CERES_DS.CAT
| - Was this file inadvertently put in this directory? Its DATA_SET_ID is not in | Looks like an extra file that should be removed. | We agree with SBN
VOLDESC.CAT, and this directory has another DS.CAT
RC | DWNC2VIR_I1A/DATA/20161007_EXTENDED_JULING/20161007_CYCLE1/VIR_VIS_1A_1_529950704_1.* | Recommend putting these two "orphan" files in the extended LAMO VIS volume instead, and document. | We agree with SBN
- These two files are VIS while every other file is IR.